# Ifremer 2026 ISMS Notes

This guide is specifically for the ISMS made for Ifremer in 2026. This model is modified slightly from prior ISMS versions and supports powering a cryotrap.

Supported firmware versions:

- Legacy-Compatible Firmware: https://github.com/Girguis-Lab/In-Situ-Mass-Spec/blob/ifremer-2026-build/Legacy-ISMS-firmware-Ifremer-2026/Legacy-ISMS-firmware-Ifremer-2026.ino

- Modern Firmware: https://github.com/Girguis-Lab/In-Situ-Mass-Spec/tree/ifremer-2026-build/

To install either firmware, see the Firmware Installation steps in the README.md file

Construction Differences from prior ISMS versions

- This ISMS uses newer vacuum components – the Pfeiffer MVP 010-3 DC roughing pump (model 2000020747) and Pfeiffer HiPace 80 Neo Turbomolecular pump (model 2000013626)  with TC80 control unit.

- Control electronics for unused features have been removed.

- Extra 12v DC-DC converter added to provide more 12v power for the Arduino and Mini PC.

- Control of the cryotrap power is used in place of PH probe power control in prior models.

- Grounding layout is different to remove noise and serial communication errors.

Optional Modifications

- The Topside serial communication channel from the onboard Arduino is converted to RS232 for longer distance communication. If only TTL serial is desired, the TTL  signals can be found on PINs 0 (Arduino pin RX0) and 1 (Arduino pin TX1) on the Arduino mega breakout board. Unplug the  RX/TX header going to the TTL to RS232 converter if using TTL only.

- As built, the RGA RS232 cable has all 9 wires connected – many can be omitted and the instrument will still function. Omitting wires can be useful for routing the RS232 directly to the surface. Only the RS232 TX and RX wires from the RGA DB9 Connector are technically needed  (assuming the below modifications are made and a ground  wire is already routed to the surface)

  - To make a minimal Rs232 cable work both connectors need to be made in a special way to trick the computer and RGA into thinking they are always connected when powered.

  - Follow the below wiring diagram to pass the minimum required number of wires between the RS232 RGA and laptop topside RS232 Adapter:

Additional details about proper wiring of the bottom side and topside connectors can be found in the isms v3 documentation here (connector on the RGA side):

https://github.com/Girguis-Lab/In-Situ-Mass-Spec/tree/ifremer-2026-build/documentation/Fabrication/Wiring#rga-serial-connection

And here (connector on the operator laptop side):

https://github.com/Girguis-Lab/In-Situ-Mass-Spec/tree/ifremer-2026-build/documentation/Fabrication/Wiring#rga-comms-serial-plug

Document Continues….

> Major component schematic shown with only ground wires visible:

Document Continues….

> Major component schematic shown with power wires visible: