# IFREMER 2026 ISMS Modern Firmware Guide

This guide is specifically for the ISMS made for IFREMER in 2026. This model is modified slightly from prior ISMS versions and supports powering a cryotrap. The ISMS firmware documented here is a modern firmware available that follows a different command response structure which is easier to use and offers better debugging than the legacy ISMS style.

## Firmware Install

The ISMS will come with the legacy-compatible firmware pre-installed, to use the modern firmware, see the README.md

## Serial Studio Dashboard

ISMS Modern firmware has a visual dashboard that utilizes the free Serial Studio software found here: https://serial-studio.com

- Download the project file from here and open it in Serial Studio.

- Choose data export options to record the system data while the instrument is running.
-> CSV Spreadsheet and Console export options are recommended.

- Select the COM port for the ISMS with the shown options.

- Hit connect in the upper right -> If the com port is correct, The ISMS will restart and you will see “Waiting for startup…” followed by more lines and the dashboard will open. If you see garbled messages, nothing, or data is not updating try a different com port, ensure serial parameters are as shown or try adding a null modem plug to the RS232 lines - RX & TX get swapped all the time.

The dashboard should look like this - if the raw console opens in a new window, you can safely close it.

At the top are action buttons for controlling the ISMS:

| In the bottom left corner are any active warnings coming from the ISMS as well as utility buttons such as to open the raw serial console, a stopwatch, a play/pause control for the incoming data, etc.… | In the lower right corner, you can choose different views of the data such as plots or the default overview. |
| --- | --- |

To access the command line interface of the ISMS, use the console button. Commands can be typed in the box at the bottom, and the output will be displayed on the console.

## Command Line Interface

Commands are case insensitive, and use underscores for the command and space between the command and argument(s)

### Standard Commands

HELP DEBUG (optional)

- Example:  HELP DEBUG 
Shows a help message; add DEBUG to show additional troubleshooting commands.

HELP DEBUG (optional)

- Example:  HELP DEBUG
Shows a help message; add DEBUG to show additional troubleshooting commands.

BEAT

- Example:  BEAT
Send BEAT to stop autostart for this boot cycle; used to temporarily enforce fully manual control.

STARTUP

- Example:  STARTUP
Run full auto startup sequence.

ROUGHING ON|OFF

- Example:  ROUGHING ON
Turns the roughing pump power on or off.

CRYO ON|OFF

- Example:  CRYO OFF
Turns the cryotrap power on or off.

FLUIDPUMP ON|OFF

- Example:  FLUIDPUMP ON
Turns the fluid pump power on or off.

FLUIDPUMP_RATE -100-100

- Example:  FLUIDPUMP_RATE -50
Sets the fluid pumping rate as a percentage of full speed. Negative values run the pump in reverse. 0 will turn off pump power.

TURBO ON|OFF

- Example:  TURBO ON
Power ON/OFF turbo pump at configured speed.

TURBO_SPEED 0.0-100.0

- Example:  TURBO_SPEED 75.0
Sets the turbo pump to run at a target speed as percent of max speed (100% is 90,000 RPM for the Pfeiffer TC80). Send 0 to reset to Pfeiffer default speed control mode.

TURBO_LIMIT_PWR 0-100

- Example:  TURBO_LIMIT_PWR 60
Sets the turbo pump power limit as a percentage of full power (Pfeiffer max draw is 5A).

TURBO_CLEAR_ERRORS

- Example:  TURBO_CLEAR_ERRORS
Clears turbo pump errors and warning messages.

STATS_INTERVAL MS|OFF

- Example:  STATS_INTERVAL 1000
Sets how often the system broadcasts stats to topside serial in milliseconds. Send OFF to disable stats logging.

AUTOSTART ON|OFF|DELAY

- Example:  AUTOSTART 5000
Set whether the system should startup automatically. If an integer is passed, it is the delay in milliseconds after boot.

### Troubleshooting Commands

VERSION

- Example:  VERSION
Prints firmware version.

PINOUT

- Example:  PINOUT
Prints configured Arduino pinout and I2C/RS485 addresses.

DEBUG_LOGGING OFF|LOW|HIGH

- Example:  DEBUG_LOGGING HIGH
Enable additional logging to serial for troubleshooting purposes.

RESET_SETTINGS

- Example:  RESET_SETTINGS
Resets settings saved on the ISMS to their default values (does not change TURBO pump or RGA parameters, which are saved within their respective instrument).

TURBO_RESET

- Example:  TURBO_RESET
Resets turbo pump parameters for default operation in normal conditions (light gases, default speed control mode, power limit 100%). Use this if the turbo is misbehaving.

TURBO_QUERY PARAMETER_NUMBER

- Example:  TURBO_QUERY 398
Query a parameter from the turbo pump with the 3-digit parameter number. See TC80 Manual Page

TURBO_CMD PARAM DATA

- Example:  TURBO_CMD 702 00123
Send a command to the turbo pump with the 3-digit parameter number and optional data string.

TURBO_RAW COMMAND

- Example:  TURBO_RAW 0010039802=?115
Send raw ASCII to the turbo pump. See

GPIO PIN_NUMBER ON|OFF

- Example:  GPIO 13 ON
Set any Arduino pin high or low – TESTING ONLY, potentially dangerous.

## Output Format

The modern code serial outputs follow a consistent pattern:

- Lines starting with  +    |     ->  or  <-   are responses from the ISMS that can be ignored by system stats parsing code.

- Lines starting with  ->  are messages sent to the Turbo Pump

- Lines starting with  <-  are responses from the Turbo Pump

- Lines starting with    !    are warnings or errors.

Other lines are the system status update that is sent periodically and follows a key-value pair format:

Fluidpump_Rate:⁠-50%,Turbo_RPM_Setpoint:90000,Turbo_RPM:32000(!),

Turbo_Current:4.40A(!),Turbo_Rotor_Temp:21C,Turbo_Elec_Temp:28C,

Turbo_Bottom_Temp:21C,Roughing_on:1,Fluidpump_on:1,Cryo_on:0

- Keys and values are separated by a colon, and each pair is separated by comma.

- Values are always numeric but may start with minus and end in a unit and/or the characters (!) which indicate that the value is outside of the expected range of normal operation. Note that the (!) is expected to appear while the instrument is pumping out or while the instrument is spinning down.

- Values for Roughing_on, Fluidpump_on, and Cryo_on indicate if that component is powered where 0 indicates OFF and 1 indicates ON