Ifremer 2026 ISMS Operation

This guide is specifically for the ISMS made for Ifremer in 2026. This model is modified slightly from prior ISMS versions and supports powering a cryotrap. The ISMS firmware documented here closely follows the previous ISMS versions command format with minor differences, there is also a more modern firmware available that follows a different command response structure which is easier to use and offers better debugging.

## Use

Outlined below are the common commands and order of operations for running the ISMS at sea (using the terminal). As you might expect, this is not an exhaustive manual for the instrument but should be able to help you accomplish the typical operation of the instrument for benchtop and in-water use. The instructions here will assume benchtop operation; the same commands and considerations apply to in-water use, but some of the diagnostic feedback (sounds, etc.) will be different and you’ll have to rely on time/data to guide your moves.

First, connect 2x USB to RS-232 converters to the laptop. The computer will assign them port numbers (often large numbers) and you can determine what the numbers are by going to the Windows ‘Device Manager’ and plugging/unplugging them to see what number is assigned to what port. Next, connect the instrument to the computer using the 2x RS-232 cables coming off the deck cable using the test cable connected to the MCIL8F connector. Ensure both RS-232 connectors are connected directly (i.e., without reversers, null modems, or the FHS (‘special’) cable).

Open 2x terminal windows (Arduino, Termite, HyperTerminal, PuTTY, etc.) ensure they are properly set up for serial communications with the Arduino (correct com port, 9600 baud, 8 data bits, 1 stop bit, no parity, no flow control). Each of the two terminal windows should correspond one of the comm assignments determined above.

Next, plug in the instrument. When you power it up, on ONE of the terminal windows, you should see the typical splash message indicating the instrument is alive. The other will show minimal, if any, feedback. The one that has the splash screen is your Arduino comm port, the second terminal window can be closed. IF NEITHER SHOW THE SPLASH SCREEN, attach a reverser/null modem to BOTH comm ports and try again; RX/TX gets swapped all the time and this is a common source of error.

With the instrument on, the data should be broadcast at 30 second intervals. If you don't see data, this indicates an issue with the comm port (or, more unlikely, an issue with the Arduino). You can cycle power on the computer or instrument to see if this resolves itself. The USB to RS-232 converters are common sources of error; swap them out, restart the computer, and try again.

Commands that you send to the instrument start with U and responses from the instrument generally start with M or LOG. For example, if you send U BEAT, the instrument will return M BEAT, or U PKNF1ON and the instrument returns M OK. If the command isn't in a format that the instrument understands or if the command fails to produce the desired result, the instrument will respond M EN. Commands should be sent in ALL CAPS.

Below is a list of common direct commands. Following this list will be a series of commands to start operating the instrument from powered down and completely vented.

## Common Direct Commands:

- U BEAT: tells the instrument that you are there and that you will be issuing manual commands

- U PROUGHON: turns on the roughing vacuum pump

- U PROUGHOFF: turns off the roughing vacuum pump

- U PTMPON: turns on the turbo board/begins looking for data from the turbo pump. This command doesn't start spinning the pump, but rather just allows you to view the state of the pump

- U PTMPOFF: turns off the turbo board/stops looking for data from the turbo pump. This command doesn't stop spinning the pump, but stops the stream of data from it

- U TON: start spinning the turbo pump to max speed

- U TOFF: stop spinning the pump via passive braking

- U TSXXX: percent of full speed operation (100 corresponds to full speed, 050 to half speed, etc.; TS100 is default). This value must have 3 digits.

- U TPXXX: percent of allowable power draw (100 corresponds to full power (~5A in Pfeiffer pumps); TP050 is default). This value must have 3 digits.

- U PCRYOON: turns on the cryo trap power

- U PCRYOOFF: turns off the cryo trap power

- U PFLUIDON: turns on the sea water/fluid pump. In the case of using the gear pump, you will also have to set the pumping rate via the below commands (U SPRATEXXX) to have the pump actually begin moving fluid

- U PFLUIDOFF: turns off the sea water/fluid pump.

- U SPRATEXXX: (applicable to gear pump only) - sets the pumping rate on the sea water/fluid pump to XXX percent of full speed. For example, U SPRATE005 is equal to 5% of full speed, U SPRATE050 is 50%, and U SPRATE100 is 100% of full speed. This value must have 3 digits. 100% is 5V and the voltage scales linearly with percent applied (i.e., 50% is 2.5V), and the approximate flow rate can be calculated using the provided spreadsheet.

- U SRATEDXXXXXXX: this changes the rate at which data is broadcast to the terminal and XXXXXXX is the amount of time (in ms) between each broadcast. Similarly, this must always have 7 digits (e.g., a broadcast interval of 5 seconds would be U SRATED0005000.

From a dead instrument, the below series of commands will start the instrument, all associated peripherals, and prepare you for in-water sampling.

Once the instrument has been powered on and you see the ‘splash’ in the comm port, send the following commands in the order shown:

- U BEAT (this tells the instrument that you’re in control, must be sent within 60 seconds of power on or autostart will begin; autostart can be killed with a power cycle)

- U SRATED0005000 (set the data broadcast rate to once every 5 seconds)

- U PROUGHON (start the roughing pump)

- Allow the roughing pump to operate for 2 mins to clear the chamber

- U PTMPON

- Ensure the turbo data are populating in the instrument data stream

- U TON

- Wait until the turbo reaches 90k RPM

- U PCRYOON (turn on the Cryotrap)

- U PFLUIDON (turn on the water/fluid pump)

- U SPRATE050 (set the fluid pump rate to 50% or whatever rate you’d like to pump).

- Start RGA software and begin scanning.

# Full command list

### Beat Command

| Command | Description | Notes |
| --- | --- | --- |
| U BEAT | Tells the instrument that the host controller is active and the ISMS should not auto run anything. | Must be sent within 60 seconds of power-on, or autostart will begin. Autostart can be stopped only by a power cycle. |

### Query Commands

| Command | Description | Notes |
| --- | --- | --- |
| U ?ID | Get system ID. |  |
| U ?FIRMWARE | Get firmware version. |  |
| U ?DATA | Output one data line on command. |  |
| U ?2C | Perform an I²C bus scan. |  |
| U ?RAM | Get current free RAM. |  |

### Power Commands

| Command | Description | Notes |
| --- | --- | --- |
| U PTMPON U PTMPOFF | Enable / disable turbo molecular pump monitoring. | Does not switch power to the turbo pump in this version. |
| U PROUGHON U PROUGHOFF | Turn roughing vacuum pump on / off. |  |
| U PRGAON U PRGAOFF | Turn RGA power on / off. | Not applicable to this model. |
| U PCRYOON U PCRYOOFF | Turn cryo cooler pump power on / off. |  |
| U PFLUIDON U PFLUIDOFF | Turn fluid pump power on / off. |  |

### Set / Configuration Commands

| Command | Description | Parameters / Notes |
| --- | --- | --- |
| U SRESETCOM1 | Reset serial communication with the turbo pump. | Serial1 is used to communicate with the turbo pump. |
| U SRATEDxxxxxxx | Set data message broadcasting interval | xxxxxxx = zero-padded 7-digit interval in milliseconds. |
| U SRATETxxxxxxx | Set turbo pump information polling interval. | xxxxxxx = zero-padded 7-digit interval in milliseconds. |
| U SDBON U SDBOFF | Enable / disable continuous data broadcasting. |  |
| U SWDON U SWDOFF | Enable / disable watchdog timer. | Disable the watchdog if unexpected restarts occur. |
| U SPRATExxx | Set fluid pump speed. | xxx = zero-padded percentage of full pumping speed (e.g., 050 = 50%). |
| U SAUTO | Start the autostart procedure immediately. |  |

### Turbo Pump Commands

| Command | Description | Parameters / Notes |
| --- | --- | --- |
| U TON U TOFF | Turn turbo pump motor on / off. | Does not control main power to the pump, only the motor command. |
| U TSxxx | Set turbo pump speed (percent of full RPM). | xxx = zero-padded percent of full RPM (e.g., 080 = 80% of maximum speed). |
| U TPxxx | Set turbo pump power limit (percent of rated power). | xxx = zero-padded percent of full power draw (rated 5A). |
| U TAE | Acknowledge / clear turbo pump error status. |  |

## Data output format

Example data string:

DATA,at1421885,TMP:3601:0:0:1470:24070:310:240:0:500:900000:000000,POWST:1:0:1:0:0:0:0:0:1:1,ENDDATA

Parts broken down:

| Label / Meaning | Raw Value | Decoded Value | Units | Notes / Description |
| --- | --- | --- | --- | --- |
| Record prefix | DATA | – | – | Every data line begins with DATA. |
| Timestamp (Arduino millis since boot) | at1421885 | 1,421,885 | ms | at followed by milliseconds since Arduino boot. |
| Section marker | TMP | – | – | Start of the turbomolecular pump data section. |
| Rotor speed | 3601 | 3,601 | RPM | Current speed of the turbo pump rotor. |
| Drive unit over-temperature flag | 0 | 0 (false) | 0/1 flag | 1 = drive unit over-temperature; 0 = normal. |
| Pump body over-temperature flag | 0 | 0 (false) | 0/1 flag | 1 = pump body over-temperature; 0 = normal. |
| Current consumption (×100) | 1470 | 1.47 | A | Current draw of the turbo pump; raw value ÷ 100. |
| Drive voltage (×100) | 24070 | 240.70 | V | Drive voltage of the turbo pump; raw value ÷ 100. |
| Electronics area temperature (×10) | 310 | 31.0 | °C | Temperature at the electronics area of the turbo; raw value ÷ 10. |
| Bottom area temperature (×10) | 240 | 24.0 | °C | Temperature at the bottom area of the turbo; raw value ÷ 10. |
| Power stage temperature (×10) | 0 | 0.0 | °C | Temperature of the power stage; raw value ÷ 10. |
| Motor rotor temperature (×10) | 500 | 50.0 | °C | Temperature of the motor rotor; raw value ÷ 10. |
| Rotor setpoint speed | 900000 | 900,000 | RPM | Target setpoint speed for the turbo pump rotor. |
| Extra field for compatibility | 000000 | – | – | Reserved/extra field to maintain compatibility with other firmware or formats. |
| Section marker | POWST | – | – | Start of the power state section. |
| Turbo | 1 | ON | 0/1 flag | Turbo pump on/off state. |
| RGA | 0 | OFF | 0/1 flag | (NOT APPLICALBLE) Residual Gas Analyzer power state. |
| Roughing Pump | 1 | ON | 0/1 flag | Roughing pump power state. |
| Roughing2 | 0 | OFF | 0/1 flag | (NOT APPLICALBLE) Roughing pump 2 power state. |
| Fan 1–4 | 0 | OFF | 0/1 flag | (NOT APPLICALBLE) Shared state for fan bank 1–4. |
| Aux24 | 0 | OFF | 0/1 flag | (NOT APPLICALBLE) Auxiliary 24 V supply state. |
| heat | 0 | OFF | 0/1 flag | (NOT APPLICALBLE) Heater state |
| aux12 | 0 | OFF | 0/1 flag | (NOT APPLICALBLE) Auxiliary 12 V rail / device state. |
| cryo | 1 | ON | 0/1 flag | Cryotrap power state. |
| fluidpump | 1 | ON | 0/1 flag | Water / fluid pump power state. |